## Sheet1
| Address [14:8] = 0 |  |  |  | 71124 |  | 71143
Register | Address [7:0] | Explanation |  | Default | AED | Line | AED | Line | Read/Write | Remark
FSM_REG | 0xff | To show Current Status (FSM) # B[2:0] : FSM status - "000" = FSM_rst - "001" = FSM_init - "010" = FSM_back_bias - "011" = FSM_flush - "100" = FSM_aed_read - "101" = FSM_exp - "110" = FSM_read # B[3] : Ready to get image - to tell a Detector ready to get image # B[4] : Aed ready done - to tell a Detector ready to get bright image (to get X-ray shot) in auto trigger mode # B[5] : expose readout exist - from expose to readout flag , active high # B[6] : panel stable exist - from back bias to flush time until saturation flus repeat count # B[15:7] : Reserved | [6 empty] | Read Only
SYS_CMD_REG | 0x01 | [15:5] | Reserved | [5 empty] | R/W
|  | [4] | system_rst - To reset FPGA Logic = Digital logic initialized | 0
|  | [3] | soft_trigger - Can be used instead of external "exp_req" signal - Keep the bit high until image data trasmission completed from FPGA to CPU | 0
|  | [2] | get_bright - '1' = getting bright images - Keep the bit high until image data trasmission completed from FPGA to CPU | 0
|  | [1] | get_dark - '1' = getting dark images - Keep the bit high until image data trasmission completed from FPGA to CPU | 0
|  | [0] | reset_FSM - '1' = FSM will be "FSM_rst" | 1
OP_MODE_REG | 0x02 | [15:9] | Reserved | [5 empty] | R/W 'h02
|  | [8] | Burst Get Image - burst image acquisition mode for dynamic mode | 0
|  | [7] | Reserved | 0
|  | [6] | en_test_roic_row '- '1' = During this bit = '1' , roic data of vsync count | 0
|  | [5] | Reserved | 0
|  | [4] | en_test_pattern_row - '1' = During this bit='1', raw test pattern generated intead of real image data | 0 | 0 | 0 | 0 | 0
|  | [3] | en_test_pattern_col - '1' = During this bit='1', column test pattern generated intead of real image data | 0 | 0 | 0 | 0 | 0
|  | [2] | en_16bit - '1' = enable 16bit ADC mode | 0 | 0 | 0 | 0 | 0
|  | [1] | en_panel_stable - '0' = Disable the mode - '1' = Back-Biasing(Saturation) operation and Flushing operation repeated in Idle state | 1 | 1 | 1 | 1 | 1
|  | [0] | en_full _flush - '0' = Disable the mode - '1' = flushing operation will be the same as Read-Out operation | 0 | 0 | 0 | 0 | 0
SET_GATE |  | [15] | stv_sel_r2 | 1 | 1 | 1 |  |  | R/W | RAON, Blue
|  | [14] | stv_sel_l2 | 1 | 1 | 1
|  | [13] | stv_sel_r1 | 0 | 0 | 0
|  | [12] | stv_sel_l1 | 0 | 0 | 0
|  | [11] | stv_sel_h | 1 | 1 | 1
|  | [10] | Reserved | 0 | 0 | 0
|  | [9] | gate_lr2 | 0 | 0 | 0
|  | [8] | gate_lr1 | 1 | 1 | 1
|  | [7] | gate_oepsn | 0 | 0 | 0
|  | [6] | gate_stv_mode | 0 | 0 | 0
|  | [5] | gate_ud | 1 | 0 | 0
|  | [4] | gate_sel | 1 | 0 | 0
|  | [3] | gate_cs2 | 1 | 1 | 1
|  | [2] | gate_cs1 | 0 | 0 | 0
|  | [1] | gate_mode2 | 1 | 0 | 0
|  | [0] | gate_mode1 | 1 | 0 | 0
GATE_SIZE | 0x04 | [15] | The number of Gate Channels - setting value : 512 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h0200 Line : 'h0200
|  | [14] |  | 0 | 0 | 0 | 0 | 0
|  | [13] |  | 0 | 0 | 0 | 0 | 0
|  | [12] |  | 0 | 0 | 0 | 0 | 0
|  | [11] |  | 0 | 0 | 0 | 0 | 0
|  | [10] |  | 0 | 0 | 0 | 0 | 0
|  | [9] |  | 1 | 1 | 1 | 1 | 1
|  | [8] |  | 0 | 0 | 0 | 0 | 0
|  | [7] |  | 0 | 0 | 0 | 0 | 0
|  | [6] |  | 0 | 0 | 0 | 0 | 0
|  | [5] |  | 0 | 0 | 0 | 0 | 0
|  | [4] |  | 0 | 0 | 0 | 0 | 0
|  | [3] |  | 0 | 0 | 0 | 0 | 0
|  | [2] |  | 0 | 0 | 0 | 0 | 0
|  | [1] |  | 0 | 0 | 0 | 0 | 0
|  | [0] |  | 0 | 0 | 0 | 0 | 0
PWR_OFF_DWM | 0x05 | [15:2] | Reserved | [5 empty] | Write Only | RAON, Blue reserved
|  | [1] | en_pwr_off - '0' = Disable Power Off Mode - '1' = Enable Power Off Mode * to Wake Up from the Power Off Mode 3sec required at least | 0 | 0 | 0 | 0 | 0 | AED : 'h0000 Line : 'h0000
|  | [0] | en_pwr_dwn - '0' = Disable Power Down Mode - '1' = Enable Power Down Mode * to Wake Up from the Power Down Mode 1sec required at least | 0 | 0 | 0 | 0 | 0
READOUT Count | 0x06 | [15:0] | Readout count | [5 empty] | Write Only | Expose - Readout 이후 readout 추가 동작 , Leg 검증용
|  |  |  | 0 | 0 | 0 | 0 | 0 | AED : 'h0000 Line : 'h0000
|  |  |  | 0 | 0 | 0 | 0 | 0
EXPOSE_SIZE | 0x10 | [15] | - Exposure time How to calculate real exposing time - Minimum Unit = 0x61A8 x "fsm_clk" clock rate - Exposing time = Minimum Unit x "Expose_size" ex) fsm_clk = 25MHz / Expose_size = 200 (0x00C8) - Minimum Unit = 0x61A8 x 40ns = 1 msec - Exposing time = 1 msec x 0x01F4 = 500 msec | 0 |  |  |  |  | R/W | Min Unit 1msec
|  | [14] |  | 0
|  | [13] |  | 0
|  | [12] |  | 0
|  | [11] |  | 0
|  | [10] |  | 0
|  | [9] |  | 0
|  | [8] |  | 1
|  | [7] |  | 1
|  | [6] |  | 1
|  | [5] |  | 1
|  | [4] |  | 1
|  | [3] |  | 0
|  | [2] |  | 1
|  | [1] |  | 0
|  | [0] |  | 0
BACK_BIAS_SIZE | 0x11 | [15] | - Back-baising period | 0 | 1 | 0 | 1 | 0 | R/W AED : 'hAFC8 Line : 'h2BD4
|  | [14] |  | 0 | 0 | 0 | 0 | 0
|  | [13] |  | 1 | 1 | 1 | 1 | 1
|  | [12] |  | 0 | 0 | 0 | 0 | 0
|  | [11] |  | 1 | 1 | 1 | 1 | 1
|  | [10] |  | 1 | 1 | 0 | 1 | 0
|  | [9] |  | 1 | 1 | 1 | 1 | 1
|  | [8] |  | 1 | 1 | 1 | 1 | 1 |  | back_bias_size : 'd45000
|  | [7] |  | 1 | 1 | 1 | 1 | 1 |  | up_back_bias : 'd2
|  | [6] |  | 0 | 1 | 1 | 1 | 1 |  | up_back_bias_opr : 'd20000
|  | [5] |  | 1 | 0 | 0 | 0 | 0 |  | down_back_bias : 'd23074
|  | [4] |  | 1 | 0 | 1 | 0 | 1 |  | down_back_bias_opr : 'd23074
|  | [3] |  | 1 | 1 | 0 | 1 | 0 |  | image_heigth : 'd3072
|  | [2] |  | 1 | 0 | 1 | 0 | 1
|  | [1] |  | 0 | 0 | 0 | 0 | 0
|  | [0] |  | 0 | 0 | 0 | 0 | 0
IMAGE_HEIGHT | 0x12 | [15] | - Image Height (the number of Gate lines) setting value : 3,072 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'hC00 Line : 'hC00
|  | [14] |  | 0 | 0 | 0 | 0 | 0
|  | [13] |  | 0 | 0 | 0 | 0 | 0
|  | [12] |  | 0 | 0 | 0 | 0 | 0
|  | [11] |  | 1 | 1 | 1 | 1 | 1
|  | [10] |  | 1 | 1 | 1 | 1 | 1
|  | [9] |  | 0 | 0 | 0 | 0 | 0
|  | [8] |  | 0 | 0 | 0 | 0 | 0
|  | [7] |  | 0 | 0 | 0 | 0 | 0
|  | [6] |  | 0 | 0 | 0 | 0 | 0
|  | [5] |  | 0 | 0 | 0 | 0 | 0
|  | [4] |  | 0 | 0 | 0 | 0 | 0
|  | [3] |  | 0 | 0 | 0 | 0 | 0
|  | [2] |  | 0 | 0 | 0 | 0 | 0
|  | [1] |  | 0 | 0 | 0 | 0 | 0
|  | [0] |  | 0 | 0 | 0 | 0 | 0
CYCLE_WIDTH_FLUSH | 0x13 | [15] | - Required Period to generate ROIC & GateIC signals for back-biasing a line | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h0064 Line : 'h0223
|  | [14] |  | 0 | 0 | 0 | 0 | 0
|  | [13] |  | 0 | 0 | 0 | 0 | 0
|  | [12] |  | 0 | 0 | 0 | 0 | 0
|  | [11] |  | 0 | 0 | 0 | 0 | 0
|  | [10] |  | 0 | 0 | 0 | 0 | 0
|  | [9] |  | 1 | 0 | 1 | 0 | 1
|  | [8] |  | 0 | 0 | 0 | 0 | 0
|  | [7] |  | 0 | 0 | 0 | 0 | 0
|  | [6] |  | 0 | 1 | 0 | 1 | 0
|  | [5] |  | 1 | 1 | 1 | 1 | 1
|  | [4] |  | 0 | 0 | 0 | 0 | 0
|  | [3] |  | 0 | 0 | 0 | 0 | 0
|  | [2] |  | 0 | 1 | 0 | 1 | 0
|  | [1] |  | 1 | 0 | 1 | 0 | 1
|  | [0] |  | 1 | 0 | 1 | 0 | 1
CYCLE_WIDTH_AED | 0x14 | - Required Period to generate ROIC & GateIC signals for AED read |  |  |  |  | 'h0E4C | 'h0E4C | R/W
CYCLE_WIDTH_READ | 0x15 | - Required Period to generate ROIC & GateIC signals for reading a line (image data) |  |  |  |  | 'h1040 | 'h1040 | R/W
REPEAT_BACK_BIAS | 0x16 | - repeat_back_bias : to define how many times repeat back-biasing - repeat_flush : to define how many times repeat flushing |  |  |  |  | 'h0001 | 'h0001 | R/W
REPEAT_FLUSH | 0x17 | [5 empty] | 'h0002 | 'h0004 | R/W
SATURATION_FLUSH_REPEAT | 0x18 | - repeat "back_bias" and "flush" FSM |  |  |  |  | 'h0002 | 'h0001 | R/W
EXP_DELAY | 0x19 | [5 empty] | 'h0064 | 'h0064 | R/W | 100
READY_DELAY | 0x1a | [5 empty] | 'h0064 | 'h0064 | R/W | 100
PRE_DELAY | 0x1b | [5 empty] | 'h0000 | 'h0000 | R/W
POST_DELAY | 0x1c | [5 empty] | 'h0000 | 'h0000 | R/W
Frame Counter | 0x1E | [7 empty] | Read Only
UP_BACK_BIAS | 0x20 | - To switch Pixel Bias between (+) bias and (-) bias for pixel saturation |  |  |  |  | 'h0002 | 'h0002 | R/W
DOWN_BACK_BIAS | 0x21 | [5 empty] | 'h5A22 | 'h19F6 | R/W
UP_BACK_BIAS_OPR | 0x22 | [5 empty] | 'h4E20 | 'h0DF6 | R/W
DOWN_BACK_BIAS_OPR | 0x23 | [5 empty] | 'h5A22 | 'h19F6 | R/W
UP_GATE_STV1_READ | 0x24 | - According to Panel condition, these register must be set up again |  |  |  |  | 'h003C | 'h003C | R/W
DOWN_GATE_STV1_READ | 0x25 | [5 empty] | 'h0168 | 'h0168 | R/W
UP_GATE_STV2_READ | 0x26 | [5 empty] | 'h003C | 'h003C | R/W
DOWN_GATE_STV2_READ | 0x27 | [5 empty] | 'h0168 | 'h0168 | R/W
UP_GATE_CPV1_READ | 0x28 | [5 empty] | 'h006E | 'h006E | R/W
DOWN_GATE_CPV1_READ | 0x29 | [5 empty] | 'h01E5 | 'h01E5 | R/W
UP_GATE_CPV2_READ | 0x2a | [5 empty] | 'h006E | 'h006E | R/W
DOWN_GATE_CPV2_READ | 0x2b | [5 empty] | 'h01E5 | 'h01E5 | R/W
DOWN_GATE_OE1_READ | 0x2c | [5 empty] | 'h03DE | 'h03DE | R/W
UP_GATE_OE1_READ | 0x2d | [5 empty] | 'h09BA | 'h09BA | R/W
DOWN_GATE_OE2_READ | 0x2e | [5 empty] | 'h03DE | 'h03D9 | R/W
UP_GATE_OE2_READ | 0x2f | [5 empty] | 'h09BA | 'h09BA | R/W
DOWN_GATE_XAO_READ | 0x30 | [5 empty] | 'h0000 | 'h0000 | R/W
UP_GATE_XAO_READ | 0x31 | [5 empty] | 'h0000 | 'h0000 | R/W
UP_GATE_STV1_AED | 0x32 | - According to Panel condition, these register must be set up again |  |  |  |  | 'h003C | 'h003C | R/W
DOWN_GATE_STV1_AED | 0x33 | [5 empty] | 'h0168 | 'h0168 | R/W
UP_GATE_STV2_AED | 0x34 | [5 empty] | 'h003C | 'h003C | R/W
DOWN_GATE_STV2_AED | 0x35 | [5 empty] | 'h0168 | 'h0168 | R/W
UP_GATE_CPV1_AED | 0x36 | [5 empty] | 'h006E | 'h006E | R/W
DOWN_GATE_CPV1_AED | 0x37 | [5 empty] | 'h01E5 | 'h01E5 | R/W
UP_GATE_CPV2_AED | 0x38 | [5 empty] | 'h006E | 'h006E | R/W
DOWN_GATE_CPV2_AED | 0x39 | [5 empty] | 'h01E5 | 'h01E5 | R/W
DOWN_GATE_OE1_AED | 0x3a | [5 empty] | 'h03DE | 'h03DE | R/W
UP_GATE_OE1_AED | 0x3b | [5 empty] | 'h09BA | 'h09BA | R/W
DOWN_GATE_OE2_AED | 0x3c | [5 empty] | 'h03DE | 'h03DE | R/W
UP_GATE_OE2_AED | 0x3d | [5 empty] | 'h09BA | 'h09BA | R/W
DOWN_GATE_XAO_AED | 0x3e | [5 empty] | 'h0000 | 'h0000 | R/W
UP_GATE_XAO_AED | 0x3f | [5 empty] | 'h0000 | 'h0000 | R/W
UP_GATE_STV1_FLUSH | 0x40 | - According to Panel condition, these register must be set up again |  |  |  |  | 'h0002 | 'h0002 | R/W
DOWN_GATE_STV1_FLUSH | 0x41 | [5 empty] | 'h0019 | 'h0025 | R/W
UP_GATE_STV2_FLUSH | 0x42 | [5 empty] | 'h0002 | 'h0002 | R/W
DOWN_GATE_STV2_FLUSH | 0x43 | [5 empty] | 'h0019 | 'h0025 | R/W
UP_GATE_CPV1_FLUSH | 0x44 | [5 empty] | 'h000A | 'h0012 | R/W
DOWN_GATE_CPV1_FLUSH | 0x45 | [5 empty] | 'h0028 | 'h0037 | R/W
UP_GATE_CPV2_FLUSH | 0x46 | [5 empty] | 'h000A | 'h0012 | R/W
DOWN_GATE_CPV2_FLUSH | 0x47 | [5 empty] | 'h0028 | 'h0037 | R/W
DOWN_GATE_OE1_FLUSH | 0x48 | [5 empty] | 'h000F | 'h001B | R/W
UP_GATE_OE1_FLUSH | 0x49 | [5 empty] | 'h0050 | 'h020F | R/W
DOWN_GATE_OE2_FLUSH | 0x4a | [5 empty] | 'h000F | 'h001B | R/W
UP_GATE_OE2_FLUSH | 0x4b | [5 empty] | 'h0050 | 'h020F | R/W
DOWN_GATE_XAO_FLUSH | 0x4c | [5 empty] | 'h0000 | 'h0000 | Reserved
UP_GATE_XAO_FLUSH | 0x4d | [5 empty] | 'h0000 | 'h0000 | Reserved
UP_ROIC_SYNC | 0x50 | [5 empty] | 'h0002 | 'h0002 | R/W
UP_ROIC_ACLK_1_READ | 0x51 | - Only 11 ACLK pulses can be generated |  |  |  |  | 'h003C | 'h003C | R/W
UP_ROIC_ACLK_2_READ | 0x52 | [5 empty] | 'h006E | 'h006E | R/W
UP_ROIC_ACLK_3_READ | 0x53 | [5 empty] | 'h0168 | 'h0168 | R/W
UP_ROIC_ACLK_4_READ | 0x54 | [5 empty] | 'h01E5 | 'h01E5 | R/W
UP_ROIC_ACLK_5_READ | 0x55 | [5 empty] | 'h03D9 | 'h03D9 | R/W
UP_ROIC_ACLK_6_READ | 0x56 | [5 empty] | 'h09C4 | 'h09C4 | R/W
UP_ROIC_ACLK_7_READ | 0x57 | [5 empty] | 'h09F6 | 'h09F6 | R/W
UP_ROIC_ACLK_8_READ | 0x58 | [5 empty] | 'h0A46 | 'h0A46 | R/W
UP_ROIC_ACLK_9_READ | 0x59 | [5 empty] | 'h0FA5 | 'h0FA5 | R/W
UP_ROIC_ACLK_10_READ | 0x5a | [5 empty] | 'h0FD7 | 'h0FD7 | R/W
UP_ROIC_ACLK_11_READ | 0x5b | [5 empty] | 'h0FDC | 'h0FDC | R/W
UP_ROIC_ACLK_1_AED | 0x5c | - Only 11 ACLK pulses can be generated |  |  |  |  | 'h003C | 'h003C | R/W
UP_ROIC_ACLK_2_AED | 0x5d | [5 empty] | 'h006E | 'h006E | R/W
UP_ROIC_ACLK_3_AED | 0x5e | [5 empty] | 'h0168 | 'h0168 | R/W
UP_ROIC_ACLK_4_AED | 0x5f | [5 empty] | 'h01E5 | 'h01E5 | R/W
UP_ROIC_ACLK_5_AED | 0x60 | [5 empty] | 'h03D9 | 'h03D9 | R/W
UP_ROIC_ACLK_6_AED | 0x61 | [5 empty] | 'h09C4 | 'h09C4 | R/W
UP_ROIC_ACLK_7_AED | 0x62 | [5 empty] | 'h09F6 | 'h09F6 | R/W
UP_ROIC_ACLK_8_AED | 0x63 | [5 empty] | 'h0A46 | 'h0A46 | R/W
UP_ROIC_ACLK_9_AED | 0x64 | [5 empty] | 'h0DB1 | 'h0DB1 | R/W
UP_ROIC_ACLK_10_AED | 0x65 | [5 empty] | 'h0DE3 | 'h0DE3 | R/W
UP_ROIC_ACLK_11_AED | 0x66 | [5 empty] | 'h0DE8 | 'h0DE8 | R/W
UP_ROIC_ACLK_1_FLUSH | 0x67 | - Only 11 ACLK pulses can be generated |  |  |  |  | 'h0005 | 'h000C | R/W
UP_ROIC_ACLK_2_FLUSH | 0x68 | [5 empty] | 'h0007 | 'h0016 | R/W
UP_ROIC_ACLK_3_FLUSH | 0x69 | [5 empty] | 'h0053 | 'h020F | R/W
UP_ROIC_ACLK_4_FLUSH | 0x6a | [5 empty] | 'h0055 | 'h0211 | R/W
UP_ROIC_ACLK_5_FLUSH | 0x6b | [5 empty] | 'h0057 | 'h0213 | R/W
UP_ROIC_ACLK_6_FLUSH | 0x6c | [5 empty] | h0059 | 'h0215 | R/W
UP_ROIC_ACLK_7_FLUSH | 0x6d | [5 empty] | 'h005B | 'h0217 | R/W
UP_ROIC_ACLK_8_FLUSH | 0x6e | [5 empty] | 'h005D | 'h0219 | R/W
UP_ROIC_ACLK_9_FLUSH | 0x6f | [5 empty] | 'h005F | 'h021B | R/W
UP_ROIC_ACLK_10_FLUS | 0x70 | [5 empty] | 'h0061 | 'h021D | R/W
UP_ROIC_ACLK_11_FLUS | 0x71 | [5 empty] | 'h0063 | 'h021F | R/W
ROIC_REG_0 | 0x72 | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h0014 Line : 'h0014
|  | [9] | Not used, keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  | [8:6] | PWR (To select power mode) - 0 = Low Power 2 mode - 1 = Low Power 1 mode - 2 = Normal mode - 3 = Unused - 4 = Sleep mode - 5 = Power-down Mode | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [5:0] | IFS (Gain) - Integer 0 ~ 63 - Integrator full-scale selection. These bits allow the user to select the appropriate full-scale range. The full-scale range is between 0.5 pC and 32 pC. | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
ROIC_REG_1 | 0x73 | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h01A8 Line : 'h01A8
|  | [9:7] | LPF - Integer 0 ~ 7 - 0 : 1.3 us - 1 : 3.9 us - 2 : 7.8 us - 3 : 11.7 us - 4 : 15.6 us - 5 : 19.5 us - 6 : 23.4 us - 7 : 31.2 us - Low-pass filter time constant. These bits allow the user to select the time constant of the low-pass filter at the output of the integrator. The time constant range is from 1.3 μs to 31.2 μs. With ADCAVG high, the time constant doubles. | 0 | 0 | 1 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
|  | [6] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  | [5] | CDS2_RESETEN - '0' = Disable - '1' = Enable, to reset CDS2 at the beginning of Line Read-Out | 1 | 1 | 1 | 1 | 1
|  | [4] | CMR_EN - '0' = Disable - '1' = Enable, to reduce common-mode errors from the AFE | 0 | 0 | 0 | 0 | 0
|  | [3] | READDOWN - '0' = Read-Out order from channel 0 to channel 255 - '1' = Read-Out order from channel 255 to channel 0 | 1 | 1 | 1 | 1 | 1 |  | 각 옵션의 Read-out 방향 변경
|  | [2] | EXTRST - '0' = resets the integrator to the REF_TFT voltage - '1' = sets REF_INT to an externally supplied reset voltage | 0 | 0 | 0 | 0 | 0
|  | [1] | ADCAVG - '0' = Disable - '1' = The results of four consecutive ADC conversions are averaged when PIPELINE_AVGEN='1' | 0 | 0 | 0 | 0 | 0
|  | [0] | Holes - '0' = integrating electrons - '1' = integrating holes | 0 | 0 | 0 | 0 | 0
ROIC_REG_2 | 0x74 | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h0007 Line : 'h000F
|  | [9] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  | [8] | PDTFTEN - '0' = disable Power-Down TFT - '1' = Keeps the input stage of the integrator biased to REF_TFT when in power-down mode | 0 | 0 | 0 | 0 | 0
|  | [7] | REFDACDIS - '0' = enable Reference DAC - '1' = Allows the user to disable the on-board reference DAC, saving 5mW | 0 | 0 | 0 | 0 | 0
|  | [6] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  | [5] | RNDOMIZE - '0' = the sampling process is performed in a sequential manner, starting with Channel 0 - '1' = random sampling to reduce fixed pattern noise | 0 | 0 | 0 | 0 | 0
|  | [4] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  | [3] | INTCLAMP - '0' = Disable - '1' = Enable internal clamp circuit | 0 | 0 | 1 | 0 | 1
|  | [2] | DOUTMODE - '0' = Serial mode. The conversion data is output serially on the DOUTAx pins. - '1' = data is output on both the DOUTAx and DOUTBx pins. Line times of <32 us are also enabled. | 1 | 1 | 1 | 1 | 1 |  | Fixed -> Optional
|  | [1] | ECHOCLK - This bit enables DCLKOx, the optional LVDS echo clock output. | 1 | 1 | 1 | 1 | 1 |  | Fixed -> Optional
|  | [0] | Pipeline - '0' = AFE sampling and the data conversion process occur serially - '1' = AFE sampling and the data conversion process occur in parallel (pipeline mode) | 1 | 1 | 1 | 1 | 1
ROIC_REG3 | 0x75 | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h0032 Line : 'h0014
|  | [9] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  | [8] | AZEN - '0' = Disable - '1' = Enable | 0 | 0 | 0 | 0 | 0
|  | [7:0] | REFDAC - Integer 0 ~ 255 - These bits determine the value for the internal reference from 0.5 V to 4.5 V - LSB size = 15.625 mV - AUTO : 50, 1.3V (50 * 15.625 + 500) - LINE : 20, 0.8V (20 * 15.625 + 500) | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 1 | 0 | 1 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 0 | 1 | 0 | 1
|  |  |  | 0 | 1 | 0 | 1 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
ROIC_REG_4 | 0x76 | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h00A2 Line : 'h00A2
|  | [9:8] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [7:4] | INTRST_C - Integer 0 ~ 11 | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [3:0] | INTRST_O - Integer 0 ~ 11 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
ROIC_REG_5 | 0x77 | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h0014 Line : 'h0014
|  | [9:8] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [7:4] | CDS1_C - Integer 0 ~ 11 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  | [3:0] | CDS1_O - Integer 0 ~ 11 | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
ROIC_REG_6 | 0x78 | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h0058 Line : 'h0058
|  | [9:8] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [7:4] | CDS2_C - Integer 0 ~ 11 | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  | [3:0] | CDS2_O - Integer 0 ~ 11 | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
ROIC_REG_7 | 0x79 | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h0037 Line : 'h0037
|  | [9:8] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [7:4] | FA_CDS1 - Integer 0 ~ 11 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
|  | [3:0] | FA_CDS2 - Integer 0 ~ 11 | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
ROIC_REG_8 | 0x7a | [15:10] | Reserved, Keep these bits 0 | [5 empty] | R/W AED : 'h0069 Line : 'h0169
|  | [9] | Not used, Keep this bit '0' | 0 | 0 | 0 | 0 | 0
|  | [8] | CUSTCLAMPEN - '0' = Disable - '1' = Enable the custom clamp timing feature | 0 | 0 | 1 | 0 | 1
|  | [7:4] | CUSTCLMP_C - Integer 0 ~ 11 | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [3:0] | CUSTCLMP_O - Integer 0 ~ 11 | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
ROIC_REG_9 | 0x7b | [15:10] | Reserved, Keep these bits 0 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h0007 Line : 'h0007
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
|  | [9:0] | Not used, Keep these bits 7 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 0 | 0 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 0 | 0 | 1 | 1
ROIC_REG_10 | 0x7c | [15:10] | Reserved, Keep these bits 0 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h0000 Line : 'h0000
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [9:1] | Not used, Keep these bits 0 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [0] | PIPELINE_AVGEN - '0' = Disable - '1' = Enables ADC averaging during pipeline mode | 0 | 0 | 0 | 0 | 0
ROIC_REG_11 | 0x7d | [15:10] | Reserved, Keep these bits 0 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h0018 Line : 'h0018
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [9:7] | Not used, Keep these bits 0 | 0 | 0 | 0 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [6] | LFSR_EN - '0' = Disable - '1' = Enables | 0 | 0 | 0 | 0 | 0
|  | [5:0] | Not used, Keep these bits 24 | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
ROIC_REG_12 | 0x7e | [15:10] | Reserved, Keep these bits 0 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h0002 Line : 'h0002
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [9] | Not used, Keep this bit 0 | 0 | 0 | 0 | 0 | 0
|  | [8] | FUNCSLEEP : Functional sleep mode. - '0' = Disable - '1' = powers up the integrators when sleep mode is in use. | 0 | 0 | 0 | 0 | 0 |  | 이름 변경
|  | [7:0] | Not used, Keep these bits 2 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
ROIC_REG_13 | 0x7f | [15:10] | Reserved, Keep these bits 0 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h0023 Line : 'h0023
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [9:0] | Not used, Keep these bits 35 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
ROIC_REG_14 | 0x80 | [15:10] | Reserved, Keep these bits 0 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h002B Line : 'h002B
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [9:0] | Not used, Keep these bits 43 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 1 | 1 | 1 | 1 | 1
ROIC_REG_15 | 0x81 | [15:10] | Reserved, Keep these bits 0 | 0 | 0 | 0 | 0 | 0 | R/W AED : 'h0008 Line : 'h0008
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  | [9:0] | Not used, Keep these bits 8 | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 1 | 1 | 1 | 1 | 1
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
|  |  |  | 0 | 0 | 0 | 0 | 0
ROIC_BURST_CYCLE | 0x90 | - ROIC Burst : ROIC Serial Data-Out of 4 pixels |  |  | 'h0046 | 'h0046 | 'h00B9 | 'h00B9 | R/W | 71124 : 645ns + 50ns (> 695ns @10ns) 71143 : 1765ns (1850ns @10ns) ACLK 연동해서 수정 필요
START_ROIC_BURST_CLK | 0x91 |  |  | 'h0001 | 'h0001 | 'h0001 | 'h0001 | 'h0001 | R/W | 1 | `
END_ROIC_BURST_CLK | 0x92 |  |  | 'h0041 | 'h0041 | 'h0041 | 'h0041 | 'h0041 | R/W | 65
SEL_ROIC_REG | 0xA0 | - To select target ROIC to read out ROIC Register set up |  |  |  |  | 'h0000 | 'h0000 | Write Only
DN_AED_GATE_XAO_0 | 0xA2 | XAO_0 signal down 위치 |  | d2 |  |  |  |  | R/W
DN_AED_GATE_XAO_1 | 0xA3 | XAO_1 signal down 위치 |  | d90 |  |  |  |  | R/W
DN_AED_GATE_XAO_2 | 0xA4 | XAO_2 signal down 위치 |  | d180 |  |  |  |  | R/W
DN_AED_GATE_XAO_3 | 0xA5 | XAO_3 signal down 위치 |  | d270 |  |  |  |  | R/W
DN_AED_GATE_XAO_4 | 0xA6 | XAO_4 signal down 위치 |  | d360 |  |  |  |  | R/W
DN_AED_GATE_XAO_5 | 0xA7 | XAO_5 signal down 위치 |  | d450 |  |  |  |  | R/W
UP_AED_GATE_XAO_0 | 0xA8 | XAO_0 signal up 위치 |  | d3206 |  |  |  |  | R/W
UP_AED_GATE_XAO_1 | 0xA9 | XAO_1 signal up 위치 |  | d3296 |  |  |  |  | R/W
UP_AED_GATE_XAO_2 | 0xAA | XAO_2 signal up 위치 |  | d3386 |  |  |  |  | R/W
UP_AED_GATE_XAO_3 | 0xAB | XAO_3 signal up 위치 |  | d3476 |  |  |  |  | R/W
UP_AED_GATE_XAO_4 | 0xAC | XAO_4 signal up 위치 |  | d3566 |  |  |  |  | R/W
UP_AED_GATE_XAO_5 | 0xAD | XAO_5 signal up 위치 |  | d3657 |  |  |  |  | R/W
READY_AED_READ | 0xB0 | [5 empty] | 'h0190 | 'h0000 | R/W
AED_TH | 0xB1 | [5 empty] | 'h0008 | 'h0000 | R/W
SEL_AED_ROIC | 0xB2 | [5 empty] | 'h0FFF | 'h0000 | R/W
NUM_TRIGGER | 0xB3 | [5 empty] | 'h0002 | 'h0000 | R/W
SEL_AED_TEST_ROIC | 0xB4 | [5 empty] | 'h0020 | 'h0000 | R/W
AED_CMD | 0xB5 | # B[2] : aed_test_mode2 - '1' = select aed_test_mode2 # B[1] : aed_test_mode1 - '1' = select aed_test_mode1 # B[0] : en_aed - '0' = disable auto trigger mode - '1' = enable auto trigger mode |  |  |  |  | 'h0001 | 'h0000 | R/W
NEGA_AED_TH | 0xB6 | [5 empty] | 'h0004 | 'h0000 | R/W
POSI_AED_TH | 0xB7 | [5 empty] | 'h0005 | 'h0000 | R/W
AED_DARK_DELAY | 0xB8 | AED prepare delay = AED_DARK_DELAY + READY_AED_READ |  |  |  |  | 'h0064 | 'h0000 | R/W
Test_Reg_18 | 0xBA | [7 empty] | Read Only
Test_Reg_19 | 0xBB | [7 empty] | Read Only
Test_Reg_1A | 0xBC | [7 empty] | Read Only
Test_Reg_1B | 0xBD | [7 empty] | Read Only
AED_DETECT_LINE_0 | 0xC0 | AED detection line 중앙 1/6 위치 | [6 empty] | R/W
AED_DETECT_LINE_1 | 0xCA | AED detection line 중앙 2/6 위치 | [6 empty] | R/W
AED_DETECT_LINE_2 | 0xCB | AED detection line 중앙 3/6 위치 | [6 empty] | R/W
AED_DETECT_LINE_3 | 0xCC | AED detection line 중앙 4/6 위치 | [6 empty] | R/W
AED_DETECT_LINE_4 | 0xCD | AED detection line 중앙 5/6 위치 | [6 empty] | R/W
AED_DETECT_LINE_5 | 0xCE | AED detection line 중앙 6/6 위치 | [6 empty] | R/W
MAX_V_COUNT | 0xD0 | CSI2 IP 생성에 따라 설정 CSI2 4 pixel mode : 3072 / 4 = 768 CSI2 2 pixel mode : 3072 / 2 = 1536 | [6 empty] | Read Only
MAX_H_COUNT | 0xD1 | 256 | [6 empty] | Read Only | TI : 'h5449 ADI : 'h4144
CSI2_WORD_COUNT | 0xD2 | CSI2 IP 생성에 따라 설정 CSI2 4 pixel mode : 512 x 4 = 2048 CSI2 2 pixel mode : 512 x 2 = 1024 | [6 empty] | Read Only
IO_DELAY_TAB | 0xDC | [7 empty] | Read Only
ROIC_VENDOR | 0xDD | [7 empty] | Read Only | TI : 'h5449 ADI : 'h4144
FPGA_VER_L | 0xDE | FPGA 생성정보 [15:0] | [6 empty] | Read Only
FPGA_VER_H | 0xDF | FPGA 생성정보 [31:16] | [6 empty] | Read Only
PURPOSE | 0xF0 | # PURPOSE - D=0x0044, E=0x0045, G=0x0047, M=0x004d, N=0x004e, S=0x0053 - D=Dental, E=EOD, G=General (for General Medical), M=Multi Frame, N=NDT, S=Security # SIZE_1 - 17 : 0x3137, 14 : 0x3134, 12 : 0x3132, 10 : 0x3130, 09 : 0x3039 # SIZE_2 - 17 : 0x3137, 14 : 0x3134, 12 : 0x3132, 10 : 0x3130, 09 : 0x3039 # MAJOR_REV - 00 : 0x3030, 01 : 0x3031 ~ 99 : 0x3939 # MINOR_REV - 00 : 0x3030, 01 : 0x3031 ~ 99 : 0x3939 | [6 empty] | Read Only | 통합 소스에 맞게 정리 필요
SIZE_1 | 0xF1 | [7 empty] | Read Only | 통합 소스에 맞게 정리 필요
SIZE_2 | 0xF2 | [7 empty] | Read Only | 통합 소스에 맞게 정리 필요
MAJOR_REV | 0xF3 | [7 empty] | Read Only
MINOR_REV | 0xF4 | [7 empty] | Read Only
Test_Reg_0 | 0xF5 | [7 empty] | Read Only
Test_Reg_1 | 0xF6 | [7 empty] | Read Only
Test_Reg_2 | 0xF7 | [7 empty] | Read Only
Test_Reg_3 | 0xF8 | [7 empty] | Read Only
Test_Reg_4 | 0xF9 | [7 empty] | Read Only
Test_Reg_5 | 0xFA | [7 empty] | Read Only
Test_Reg_6 | 0xFB | [7 empty] | Read Only
Test_Reg_7 | 0xFC | [7 empty] | Read Only
Test_Reg_8 | 0xFD | [7 empty] | Read Only
Test_Reg_9 | 0xFE | [7 empty] | Read Only
fsm_back_bias
fsm_flush
repeat_back_bias
repeat_flush
fsm_back_bias
fsm_flush
Saturation_flush_repeat
row_cnt
X = Back_bias_size
cycle_width_bb
back_bias
back_bias_opr
Photo Diode Bias
(-)
(+)
(-)
Image_height
Delay_A
Delay_B
1
2
3
X
UP_BACK_BIAS
DOWN_BACK_BIAS
UP_BACK_BIAS_OPR
DOWN_BACK_BIAS_OPR
A
B
C
E
F
col_cnt
Gate_STV1 or STV2_READ
D
M = cycle_width_read
1
2
M
Gate_CPV1 or CPV2_READ
Gate_OE1 or OE2_READ
Up_Gate_STV1 or STV2_READ
Down_Gate_STV1 or STV2_READ
Down_Gate_CPV1 or CPV2_READ
Up_Gate_CPV1 or CPV2_READ
Up_Gate_OE1 or OE2_READ
Down_Gate_OE1 or OE2_READ
Gate_XAO_READ
Up_Gate_XAO_READ
Down_Gate_XAO_READ
A
B
C
E
F
col_cnt
Gate_STV1 or STV2_AED
D
M = cycle_width_aed
1
2
M
Gate_CPV1 or CPV2_AED
Gate_OE1 or OE2_AED
Up_Gate_STV1 or STV2_AED
Down_Gate_STV1 or STV2_AED
Down_Gate_CPV1 or CPV2_AED
Up_Gate_CPV1 or CPV2_AED
Up_Gate_OE1 or OE2_AED
Down_Gate_OE1 or OE2_AED
Gate_XAO_AED
Up_Gate_XAO_AED
Down_Gate_XAO_AED
A
B
C
E
F
col_cnt
Gate_STV1 or STV2_FLUSH
D
M = cycle_width_flush
1
2
M
Gate_CPV1 or CPV2_FLUSH
Gate_OE1 or OE2_FLUSH
Up_Gate_STV1 or STV2_FLUSH
Down_Gate_STV1 or STV2_FLUSH
Down_Gate_CPV1 or CPV2_FLUSH
Up_Gate_CPV1 or CPV2_FLUSH
Up_Gate_OE1 or OE2_FLUSH
Down_Gate_OE1 or OE2_FLUSH
Gate_XAO_FLUSH
Up_Gate_XAO_FLUSH
Down_Gate_XAO_FLUSH
A
col_cnt
ROIC_SYNC
M = cycle_width_flush
1
2
M
UP_ROIC_SYNC
A
B
col_cnt
ROIC_SYNC
B
M = cycle_width_read
1
2
M
UP_ROIC_ACLK_1_READ
UP_ROIC_ACLK_2_READ
UP_ROIC_ACLK_11_READ
ROIC_ACLK
A
B
col_cnt
ROIC_SYNC
B
M = cycle_width_aed
1
2
M
UP_ROIC_ACLK_1_AED
UP_ROIC_ACLK_2_AED
UP_ROIC_ACLK_11_AED
ROIC_ACLK
A
B
col_cnt
ROIC_SYNC
B
M = cycle_width_flsuh
1
2
M
UP_ROIC_ACLK_1_FLUSH
UP_ROIC_ACLK_2_FLUSH
UP_ROIC_ACLK_11_FLUSH
ROIC_ACLK
X
-
.
-
Purpose
Major Rev
Minor Rev
Panel Size
SIZE_1
SIZE_2
1
2
3
4
M
1
2
3
4
M
M = ROIC_BURST_CYCLE
64 bits serial data
64 bits serial data
START_ROIC_BURST_CLK
END_ROIC_BURST_CLK
ROIC burst clock
row_cnt
X = Back_bias_size
cycle_width_bb
back_bias
back_bias_opr
Photo Diode Bias
(-)
(+)
(-)
Image_height
Delay_A
Delay_B
1
2
3
X
Delay_A : Time to be (+)Bias (Refer to switch IC SPEC)
Delay_B : Time to be (-)Bias (Refer to switch IC SPEC)
How to calculate Delay_A and Delay_B
- Minimum Unit = cycle_width_bb x "fsm_clk" clock rate 4160 x 25MHz(40ns) = 166400
- Delay_A = Minimum Unit x (up_back_bias - up_back_bias_opr) setting value : 19998
- Delay_B = Minimum Unit x (back_bias_size - down_back_bias) setting balue : 21926
col_cnt
Gate_CPV_BB
roic_rst
Gate_oe_BB
N = cycle_width_bb
1
2
3
N